Metrix GX1010 Function Generator

________________________________________________________________________________

Introduction:

This instrument driver provides programming support for Metrix GX1010 Function Generator.

It contains functions for opening, configuring, taking measurements from, and

closing the instrument.

________________________________________________________________________________

Assumptions:

To successfully use this module, the following conditions must be met:

For GPIB instrument drivers:

-  the instrument is connected to the GPIB.

-  the GPIB address supplied to the initialize function must match the

GPIB address of the instrument.

For VXI instrument drivers:

-  the instrument is installed in the VXI mainframe and you are

using one of the following controller options:

Embedded controller

MXI

MXI2

GPIB-VXI

-  the logical address supplied to the initialize function must

match the logical address of the instrument.

For RS-232 instrument drivers:

-  the instrument is connected to the RS-232 interface.

-  the COM port, baud rate, parity, and timeout supplied to the

initialize function must match the settings of the instrument.

________________________________________________________________________________

Error and Status Information:

Error codes and warnings are returned as the return value of each instrument

driver function. A program should examine this value after each call to an

instrument driver function to determine if any error occurred.  The general

meaning of the return value is as follows:

Value                   Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

Possible error codes and their meanings are listed with the corresponding

instrument driver function.

________________________________________________________________________________

How To Use This Document:

This document is intended to be used as a programming reference manual.

It describes each function in the

Metrix GX1010 Function Generator

instrument. The functions are listed in alphabetical order with a

description of the function, C syntax of the function, a description of each

parameter, and possible error codes.

________________________________________________________________________________

Function Tree Layout:

Metrix GX1010 Function Generator       Function Name:

Initialize                                  gx1010_init

Configure

Main Generator                           gx1010_config_main

gx1010_main_offset

gx1010_main_symmetry

gx1010_main_phase

gx1010_main_zout

Modulation

AM and VCA                            gx1010_config_am

Sweep                                 gx1010_config_sweep

FSK                                   gx1010_fsk

Hop                                   gx1010_hop

Trigger (Burst)                          gx1010_trig

Gate                                     gx1010_gate

Set Status                               gx1010_set_status

Send Trigger                                gx1010_send_trg

Local Storage                               gx1010_loc_stores

Get Status                                  gx1010_get_status

Miscellaneous

Send command string                      gx1010_snd_cmd

Get Response string                      gx1010_get_resp

File operations                          gx1010_file_ops

Arbitary

New ARB/STAIRCASE                     gx1010_new_arb

ARB Local storage                     gx1010_arb_stores

Close                                       gx1010_close

________________________________________________________________________________

Metrix GX1010 Function Generator

Instrument Name: Metrix GX1010 10MHz

DDS Function Generator.

Description:     This instrument module provides

programming support for the

GX1010.  The module is divided

into the following functions:

Functions/Classes:

(1) Initialize

Initialize the instrument and optionally

interrogate the ID and set to the specified

reset conditions.

(2) Configure...

This class of functions configures the main

generator and allows access

to the Modulation capabilities of the GX1010.

(1) Main Generator

Set up such GX1010 Main Generator

parameters as  output function,

frequency and amplitude.

(2) Modulation...

This class of functions configures and

controls the Modulation capabilities

of the GX1010.

(1) AM and VCA

Amplitude Modulation.

(2) Sweep

Software controlled Frequency.

(3) FSK

Frequency shift keying

(4) Hop

Set up and run Hop mode

(3) Trigger (Burst)

Set up and control the Trigger

capabilities of the GX1010.

(4) Gate

Set up and control the Trigger

capabilities of the GX1010.

(5) Set Status

Configure the IEEE 488.2 status

reporting capabilities of the GX1010.

(3) Local Storage

Executes the Store and Recall commands.

(4) Get Status

Returns the current value of the specified

IEEE 488.2 status byte.

(5) Miscellaneous

This class of functions allows access to

additional GX1010 remote control commands

which is not available elsewhere in this

driver. Also included are file operations.

(1) Send Command String

(2) Get Response String

(3) File Operations

(4) Arbitary

This class of functions gives access

to the arbitary waveform capabilities

of the GX1010

(1) New ARB

(2) ARb Local storage

(3) Define Staircase waveform

(6) Close

Take the instrument offline.

________________________________________________________________________________

The following functions are in alphabetical order.

________________________________________________________________________________

gx1010_arb_stores

int gx1010_arb_stores (int instrumentID, int operation, int storeNumber,

char name[]);

Purpose

Panel Name:    ARB Local storage

Description:   Allows ARB waveforms to be stored

and recalled from local

non-volatile memory. The current

ARB is always the last recalled or

downloaded.

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   Accepts the Instrument ID returned

by the Initialize function to

select the desired instrument.

Variable Type: Integer

Valid Range:   1 to  Maximum Instances of

this Instrument

The Maximum Instances of this

Instrument is the maximum number

of instruments that can be

controlled by this driver in a

single system. The value is set

internally in the driver. The

typical value is 10.

operation

Variable Type       int

Control Name:  Operation

Description:   Selects the recall or store

operation.

Variable Type: Integer

Valid Range:   0 = Recall

1 = Store

Default Value: 0 - Recall

storeNumber

Variable Type       int

Control name:   Store Number

Description:    Select the number of the store to

Recall from or Store to.

Variable type:  integer

Valid range:    1 to max number of stores

Default value:  1

Note: the max number of stores will depend upon

the software level of the instrument V1.1 has 21.

However only the first 5 are writeable and an

error will result if writing is attempted above

this number.

name

Variable Type       char []

Control name:   Name

Description:    The name to use for a store

operatrion.

Variable type:  string - max length 16 characters

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   instrument ID out of range

-2   Operation value out of range

-3   Store number illegal

228   Error writing to instrument from file

229   Error reading from instrument to file

232   Instrument not initialized

240   Instrument timed out

________________________________________________________________________________

gx1010_close

int gx1010_close (int instrumentID);

Purpose

Panel Name:    Close

Description:   Closes the instrument and takes it

off line.

Note: The  instrument must  be reinitialized

to use it again.

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   Accepts the Instrument ID returned

by the Initialize function to

select the desired instrument.

Variable Type: Integer

Valid Range:   1 to  Maximum Instances of

this Instrument

The Maximum Instances of this

Instrument is the maximum number

of instruments that can be

controlled by this driver in a

single system. The value is set

internally in the driver. The

typical value is 10.

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Instrument ID out of range

221   Unable to close instrument

232   Instrument not initialized

________________________________________________________________________________

gx1010_config_am

int gx1010_config_am (int instrumentID, int enable, int waveform, int modDepth,

double modFrequencyHz);

Purpose

Panel Name:    AM and VCA

Description:   Enables AM and sets up Waveform,

Modulation Depth and Frequency.

If the modulation source is

external VCA mode is selected.

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   Accepts the Instrument ID returned

by the Initialize function to

select the desired instrument.

Variable Type: Integer

Valid Range:   1 to  Maximum Instances of

this Instrument

The Maximum Instances of this

Instrument is the maximum number

of instruments that can be

controlled by this driver in a

single system. The value is set

internally in the driver. The

typical value is 10.

enable

Variable Type       int

Control Name:  Enable

Description:   Sets the status of AM.

Variable Type: Integer

Valid Range:

0 = Off

1 = On

2 = On and configure

Default Value: 0 - Off

Note: If set to ON or OFF all other controls

are ignored, to use these you must use

On + config.

waveform

Variable Type       int

Control Name:  Waveform

Description:   Selects the modulation waveform.

Variable Type: Integer

Valid Range:   0 = Sine  (1kHz)

1 = Square

2 = External (VCA mode)

Default Value: 1

Note: If sinewave modulation is selected The

Modulation frequency (TGEN) will be fixed to

1kHz. Otherwise the TGEN will be left free for

other modes. Selecting sinewave causes the Mod

frequency control to be ignored.

modDepth

Variable Type       int

Control Name:  Mod Depth

Description:   Sets the Depth of modulation

expressed as a percentage of the

carrier. This parameter has no

effect if the modulation Waveform

is External.

Variable Type: integer

Valid Range:   0% to 100%

Default Value: 30

modFrequencyHz

Variable Type       double

Control Name:  Mod Frequency

Description:   Sets the internal modulation

frequency in Hz.This parameter has

no effect if the modulation

Waveform is External.

Note. This parameter is actually

the internal Trigger generator

frequency.

Variable Type: double

Valid Range:   5e-3 to 50e+3 (5mHz to 50kHz)

Default Value: 1000

Note: The internal trigger generator period has

a setting resolution of 20us max. This

driver converts the requested frequency to a

period value and sends the resulting number to

the GX1010. This may cause a warning that the

20us resolution has been exceeded. This can

usually be ignored.

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   instrument ID out of range

-2   Enable value out of range

-3   Waveform value out of range

-4   Mod depth value out of range

-5   Frequency value out of range

230   Error writing to instrument

232   Instrument not initialized

240   Instrument timed out

________________________________________________________________________________

gx1010_config_main

int gx1010_config_main (int instrumentID, int function, int output,

double frequencyHz, double pkPkOutputV, int sense);

Purpose

Panel Name:    Main Generator

Description:   Sets up Function, Frequency,

Peak-Peak Output level and Output

status for the Main Generator.

Secondary functions are provided

to set DC Offset, Symmetry, Phase

and output impedance.

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   Accepts the Instrument ID returned

by the Initialize function to

select the desired instrument.

Variable Type: Integer

Valid Range:   1 to  Maximum Instances of

this Instrument

The Maximum Instances of this

Instrument is the maximum number

of instruments that can be

controlled by this driver in a

single system. The value is set

internally in the driver. The

typical value is 10.

function

Variable Type       int

Control Name:  Function

Description:   Specifies the output waveform

function.

Variable Type: Integer

Valid Range:   0 = Sine

1 = Triangle

2 = Square

3 = Positive Pulse

4 = Negative Pulse

5 = Positive Ramp

6 = Negative Ramp

7 = Staircase

8 = ARB

9 = Noise

Default Value: 0 - Sine

output

Variable Type       int

Control Name:  Output

Description:   Turns the Main output on or off.

Variable Type: Integer

Valid Range:   1 = On

0 = Off

Default Value: 1 - On

frequencyHz

Variable Type       double

Control Name:  Frequency

Description:   The output Frequency expressed in

Hz.

Variable Type: double

Valid Range:   1e-4 to 10e+6 (0.1mHz to 10MHz)

Default Value: 10kHz

pkPkOutputV

Variable Type       double

Control Name:  Pk-Pk Output

Description:   Specifies the output level as a

peak to peak EMF value expressed

in Volts.

Variable Type: double

Valid Range:   Waveform function dependent. See

GX1010 instruction manual for more

information.

Default Value: 20V

sense

Variable Type       int

Control Name:  Sense

Description:   Sets the output sense to NORMAL or

INVERTed.

Variable Type: Integer

Valid Range:   1 = Invert

0 = Normal

Default Value: 0

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   instrument ID out of range

-2   Function value out of range

-3   Output value out of range

-4   Frequency value out of range

-5   DC offset value out of range

-6   Symmetry value out of range

-7   Phase value out of range

-8   Zout value out of range

-9   Sense value out of range

230   Error writing to instrument

232   Instrument not initialized

240   Instrument timed out

________________________________________________________________________________

gx1010_config_sweep

int gx1010_config_sweep (int instrumentID, int enable, int mode, int law,

double timesec, double startfrequencyHz,

double stopfrequencyHz, int triggerSource,

double markerB, double markerA);

Purpose

Panel Name:    Sweep

Description:   Enables Sweep (Software Controlled

Frequency) and sets up the Mode,

Start and Stop Frequencies and the

sweep time.

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   Accepts the Instrument ID returned

by the Initialize function to

select the desired instrument.

Variable Type: Integer

Valid Range:   1 to  Maximum Instances of

this Instrument

The Maximum Instances of this

Instrument is the maximum number

of instruments that can be

controlled by this driver in a

single system. The value is set

internally in the driver. The

typical value is 10.

enable

Variable Type       int

Control Name:  Enable

Description:   Sets the status of Sweep.

Variable Type: Integer

Valid Range:

0 = Off

1 = On

2 = On and configure

Default Value: 0 - Off

Note: If set to ON or OFF all other controls

are ignored, to use these you must use

On + config.

mode

Variable Type       int

Control Name:  Mode

Description:   Sets the Sweep mode.

Variable Type: Integer

Valid Range:   1 = Start - Stop

0 = Stop - Start

Default Value: 1 - Start - Stop

Note: Setting this control to 1 will cause the

sweep to go from the Start frequency to the Stop

frequency. 0 will reverse the direction.

law

Variable Type       int

Control Name:  Law

Description:   Sets the Sweep Law.

Variable Type: Integer

Valid Range:   0 = Linear

1 = Logarithmic

Default Value: 0 - linear

timesec

Variable Type       double

Control Name:  Time

Description:   Specifies the internal Sweep ramp

time in seconds.

Variable Type: double

Valid Range:   10e-3 to 999

Default Value: none

startfrequencyHz

Variable Type       double

Control Name:  Start Frequency

Description:   The Sweep Start Frequency

expressed in Hz.

Variable Type: double

Valid Range:   1e-4 to 10e+6 (0.1mHz to 10MHz)

Default Value: none

stopfrequencyHz

Variable Type       double

Control Name:  Stop Frequency

Description:   The Sweep Stop Frequency

expressed in Hz.

Variable Type: double

Valid Range:   1e-4 to 10e+6 (0.1mHz to 10MHz)

Default Value: none

triggerSource

Variable Type       int

Control Name:  Trigger Source

Description:   Sets the Sweep Trigger source.

Variable Type: Integer

Valid Range:   0 = Continuous

1 = Man/Remote

2 = External

Default Value: 0 - Continuous

markerB

Variable Type       double

Control Name:  Marker B

Description:   The Frequency for Marker B

expressed in Hz.

Variable Type: double

Valid Range:   1e-4 to 10e+6 (0.1mHz to 10MHz)

Default Value: none

Note: This is the Main generator frequency.

markerA

Variable Type       double

Control Name:  Marker A

Description:   The Frequency for Marker A

expressed in Hz.

Variable Type: double

Valid Range:   1e-4 to 10e+6 (0.1mHz to 10MHz)

Default Value: none

Note: This is the Marker frequency set on the

sweep menu.

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   instrument ID out of range

-2   Enable value out of range

-3   Mode value out of range

-4   Law value out of range

-5   Time value out of range

-6   Start frequency value out of range

-7   Stop frequency value out of range

-8   Trigger source value out of range

-9   Marker A frequency value out of range

-10   Marker B frequency value out of range

230   Error writing to instrument

232   Instrument not initialized

240   Instrument timed out

________________________________________________________________________________

gx1010_file_ops

int gx1010_file_ops (int instrumentID, int operation, char fileName[]);

Purpose

Panel name:     File Operations

Description:    Send or Receive a binary block to

or from the GX1010. The data will

be read from or written to a file.

Note            This operation is usefull for

LRN blocks.

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   Accepts the Instrument ID returned

by the Initialize function to

select the desired instrument.

Variable Type: Integer

Valid Range:   1 to  Maximum Instances of

this Instrument

The Maximum Instances of this

Instrument is the maximum number

of instruments that can be

controlled by this driver in a

single system. The value is set

internally in the driver. The

typical value is 10.

operation

Variable Type       int

Control Name:  Operation

Description:   Selects to send a file to the

GX1010 or Get a file from GX1010.

Variable Type: Integer

Valid Range:   0 = Send file to GX1010

1 = Get file from GX1010

Default Value: 0 - Send file

fileName

Variable Type       char []

Control name:   File name

Description:    This control specifies file name

from which the binary block will

be sent to the GX1010.

Variable type:  String - max length 50 characters

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   instrument ID out of range

-2   Operation value out of range

228   Error writing to instrument from file

229   Error reading from instrument to file

232   Instrument not initialized

240   Instrument timed out

________________________________________________________________________________

gx1010_fsk

int gx1010_fsk (int instrumentID, int enable, int triggerSource,

double frequencyA, double frequencyB, double internalGenerator);

Purpose

Panel Name:    FSK

Description:   Enables FSK and sets up parameters

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   Accepts the Instrument ID returned

by the Initialize function to

select the desired instrument.

Variable Type: Integer

Valid Range:   1 to  Maximum Instances of

this Instrument

The Maximum Instances of this

Instrument is the maximum number

of instruments that can be

controlled by this driver in a

single system. The value is set

internally in the driver. The

typical value is 10.

enable

Variable Type       int

Control Name:  Enable

Description:   Sets the status of FSK.

Variable Type: Integer

Valid Range:

0 = Off

1 = On

2 = On and configure

Default Value: 0 - Off

Note: If set to ON or OFF all other controls

are ignored, to use these you must use

On + config.

triggerSource

Variable Type       int

Control Name:  Trigger Source

Description:   Sets the FSK Trigger source.

Variable Type: Integer

Valid Range:   0 = Internal

1 = Man/Remote

2 = External

Default Value: 0 - Internal

frequencyA

Variable Type       double

Control Name:  Frequency A

Description:   FSK frequency A, actually

the main generator frequency,

expressed in Hz.

Variable Type: double

Valid Range:   1e-4 to 10e+6 (0.1mHz to 10MHz)

Default Value: none

frequencyB

Variable Type       double

Control Name:  Stop Frequency

Description:   FSK frequency B expressed in Hz.

Variable Type: double

Valid Range:   1e-4 to 10e+6 (0.1mHz to 10MHz)

Default Value: none

internalGenerator

Variable Type       double

Control Name:  Internal Generator

Description:   Sets the internal trigger

frequency in Hz.

Note. This parameter is actually

the internal Trigger generator

frequency (TGEN).

Variable Type: double

Valid Range:   5e-3 to 50e+3 (5mHz to 50kHz)

Default Value: 1000

Note: The internal trigger generator period has

a setting resolution of 20us max. This

driver converts the requested frequency to a

period value and sends the resulting number to

the GX1010. This may cause a warning that the

20us resolution has been exceeded. This can

usually be ignored.

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   instrument ID out of range

-2   Enable value out of range

-3   Trigger Source value out of range

-4   FSK frequency A value out of range

-5   FSK frequency B value out of range

-6   Internal generator value out of range

230   Error writing to instrument

232   Instrument not initialized

240   Instrument timed out

________________________________________________________________________________

gx1010_gate

int gx1010_gate (int instrumentID, int enable, int triggerSource,

double internalGenerator);

Purpose

Panel Name:    Gate

Description:   This panel allows access to the

Gating function of the GX1010

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   Accepts the Instrument ID returned

by the Initialize function to

select the desired instrument.

Variable Type: Integer

Valid Range:   1 to  Maximum Instances of

this Instrument

The Maximum Instances of this

Instrument is the maximum number

of instruments that can be

controlled by this driver in a

single system. The value is set

internally in the driver. The

typical value is 10.

enable

Variable Type       int

Control Name:  Enable

Description:   Sets the status of Gate.

Variable Type: Integer

Valid Range:

0 = Off

1 = On

2 = On and configure

Default Value: 0 - Off

Note: If set to ON or OFF all other controls

are ignored, to use these you must use

On + config.

triggerSource

Variable Type       int

Control Name:  Trigger Source

Description:   Sets the Gating trigger source.

Variable Type: Integer

Valid Range:   0 = Internal

1 = Man/Remote

2 = External

Default Value: 0 - Internal

internalGenerator

Variable Type       double

Control Name:  Internal Generator

Description:   Sets the internal trigger

frequency in Hz.

Note. This parameter is actually

the internal Trigger generator

frequency (TGEN).

Variable Type: double

Valid Range:   5e-3 to 50e+3 (5mHz to 50kHz)

Default Value: 1000

Note: The internal trigger generator period has

a setting resolution of 20us max. This

driver converts the requested frequency to a

period value and sends the resulting number to

the GX1010. This may cause a warning that the

20us resolution has been exceeded. This can

usually be ignored.

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   instrument ID out of range

-2   Enable value out of range

-3   Trigger source value out of range

-4   Internal generator value out of range

230   Error writing to instrument

232   Instrument not initialized

240   Instrument timed out

________________________________________________________________________________

gx1010_get_resp

int gx1010_get_resp (int instrumentID, char response[], int *numberOfBytes);

Purpose

Panel name:     Get response string

Description:    Reads a response from the GX1010

into a string variable.

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   Accepts the Instrument ID returned

by the Initialize function to

select the desired instrument.

Variable Type: Integer

Valid Range:   1 to  Maximum Instances of

this Instrument

The Maximum Instances of this

Instrument is the maximum number

of instruments that can be

controlled by this driver in a

single system. The value is set

internally in the driver. The

typical value is 10.

response

Variable Type       char []

Control name:   Response Stirng

Description:    This control specifies a string

variable to receive the response

from the GX1010. The string must

be long enough to receive the

response to avoid overflow. To aid

in checking this the number of

bytes put into the string is

returned from the function. The

absolute maximum length allowed is

2000 bytes.

Variable type:  String

numberOfBytes

Variable Type       int (passed by reference)

Control name:   Number of Bytes

Description:    This value is the number of bytes

returned by the GX1010 and

placed in the string.

Variable type:  Integer

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   instrument ID out of range

231   Error reading from instrument

232   Instrument not initialized

240   Instrument timed out

________________________________________________________________________________

gx1010_get_status

int gx1010_get_status (int instrumentID, int readStatus, int *value);

Purpose

Panel name:     Read Status

Description:    Returns the byte value from the

selected status register.

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   Accepts the Instrument ID returned

by the Initialize function to

select the desired instrument.

Variable Type: Integer

Valid Range:   1 to  Maximum Instances of

this Instrument

The Maximum Instances of this

Instrument is the maximum number

of instruments that can be

controlled by this driver in a

single system. The value is set

internally in the driver. The

typical value is 10.

readStatus

Variable Type       int

Control name:   Read Status

Description:    Used to select which status

register to read

Variable type:  integer

Valid range:    0 = Status Byte

1 = Standard Event Status

2 = SRQ enable

3 = Standard Event Enable

4 = Parallel Poll Enable

5 = Execution Error

Default value   0 - Status Byte

value

Variable Type       int (passed by reference)

Control name:   Value

Description:    Returns the value read from the

selected register.

Variable type:  integer

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   instrument ID out of range

-2   Status Enable value out of range

-3   Register Value out of range

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

236   Error interpreting instrument response

240   Instrument timed out

________________________________________________________________________________

gx1010_hop

int gx1010_hop (int instrumentID, int enable, int stepNumber, int function,

double frequencyHz, double pkPkOutputV, double DCOffsetV,

int lastStep, double timesec);

Purpose

Panel Name:    Hop

Description:   Enables HOP or sets up

Parameters for each step

Note: To avoid possible bus lockouts you

should not execute any other panels or functions

from this driver (except the Send Trigger panel)

while Hop is running. Always set enable to OFF

and execute this pannel first.

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   Accepts the Instrument ID returned

by the Initialize function to

select the desired instrument.

Variable Type: Integer

Valid Range:   1 to  Maximum Instances of

this Instrument

The Maximum Instances of this

Instrument is the maximum number

of instruments that can be

controlled by this driver in a

single system. The value is set

internally in the driver. The

typical value is 10.

enable

Variable Type       int

Control Name:  Enable

Description:   Sets the status/configuration of

Hop.

Variable Type: Integer

Valid Range:

0 = Off

1 = On

2 = Configure step

Default Value: 2 - Configure step

Note: If set to OFF all other controls

are ignored, RUN uses only the Last step control,

Config step uses all other controls.

Note: To avoid possible bus lockouts you

should not execute any other panels or functions

from this driver (except the Send Trigger panel)

while Hop is running. Always set enable to OFF

and execute this pannel first.

stepNumber

Variable Type       int

Control Name:  Step number

Description:   The step to configure

Variable Type: Integer

Valid Range:   0 to 15

Default Value: 0

function

Variable Type       int

Control Name:  Function

Description:   Specifies the output waveform

function.

Variable Type: Integer

Valid Range:   0 = Sine

1 = Triangle

2 = Square

3 = Positive Pulse

4 = Negative Pulse

5 = Positive Ramp

6 = Negative Ramp

7 = Staircase

8 = ARB

Default Value: 0 - Sine

frequencyHz

Variable Type       double

Control Name:  Frequency

Description:   The output Frequency expressed in

Hz.

Variable Type: double

Valid Range:   1e-4 to 10e+6 (0.1mHz to 10MHz)

Default Value: 1kHz

pkPkOutputV

Variable Type       double

Control Name:  Pk-Pk Output

Description:   Specifies the output level as a

peak to peak EMF value expressed

in Volts.

Variable Type: double

Valid Range:   Waveform function dependent. See

GX1010 instruction manual for more

information.

Default Value: 20V

DCOffsetV

Variable Type       double

Control Name:  DC Offset

Description:   Sets the output Offset voltage as

an EMF value expressed in Volts

Variable Type: double

Valid Range:   -10 to +10

Default Value: 0.0

lastStep

Variable Type       int

Control Name:  Last Step

Description:   The last step in the sequence to

execute.

Variable Type: Integer

Valid Range:   0 to 15

Default Value: 1

timesec

Variable Type       double

Control Name:  Time

Description:   The time for this step in seconds.

0 and 1ms have special meanings,

see below

Variable Type: Double

Valid Range:   0ms = Man/Remote

1ms = External

2ms to 65s = time duration

Default Value: 1 sec

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   instrument ID out of range

-2   Enable value out of range

-3   Step value out of range

-4   frequency value out of range

-5   DC offset value out of range

-6   Last step value out of range

-7   time value out of range

-8   Function vaqlue out of range

230   Error writing to instrument

232   Instrument not initialized

240   Instrument timed out

Note: To avoid possible bus lockouts you

should not execute any other panels or functions

from this driver (except the Send Trigger panel)

while Hop is running. Always set enable to OFF

and execute this pannel first.

________________________________________________________________________________

gx1010_init

int gx1010_init (int address, int IDQuery, int resetDevice, int *instrumentID);

Purpose

Panel Name:    Initialize

Description:   Initializes the instrument in the

following ways:

- opens the instrument and sets

its GPIB address to the speci-

fied value

- optionally performs an

identification query on the

instrument

- optionally sets the instrument

to a default configuration

- returns an instrument ID which

is used to identify the

instrument in subsequent calls.

Parameter List

address

Variable Type       int

Control Name:  Address

Description:   Specifies the address on the

GPIB to be used to communicate

with the instrument.  This value

must correspond to the current

internal setting of the

instrument.

Variable Type: Integer

Valid Range:   0 to 30

Default Value: 5

IDQuery

Variable Type       int

Control Name:  ID Query

Description:   Specifies if an ID Query is sent

to the instrument during the

initialization procedure.

Variable Type: Integer

Valid Range:   1 = Yes

0 = No

Default Value: 1 - Yes

NOTE:  Under normal circumstances the ID Query

insures that the instrument initialized

over the GPIB is the type supported by

this driver. However circumstances may

arise where it is undesirable to send an

ID Query to the instrument. In those

cases; set this control to Skip Query

and this function will initialize the

GPIB and the Command arrays in the driver,

without doing an ID Query.

resetDevice

Variable Type       int

Control Name:  Reset

Description:   Specifies if the instrument is to

be reset during the initialization

procedure.

Variable Type: Integer

Valid Range:   0 = No Reset

1 = Reset device

Default Value: 1 = Reset

NOTE: If you do not want the instrument reset.

Set this control to No Reset while

initializing the instrument.

instrumentID

Variable Type       int (passed by reference)

Control Name:  Instrument ID

Description:   Returns an Instrument ID that is

used in all subsequent function

calls to select the instrument at

the Address specified.

If more than one instrument of the

same type is used, this value will

be used to differentiate between

them. The first instrument

initialized will be assigned an

Instrument ID of 1.

Variable Type: Integer

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   GPIB Address out of range

-2   ID Query out of range

-3   Reset Device out of range

220   Unable to open instrument

221   Unable to close instrument

223   Instrument identification query failed

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

240   Instrument timed out

________________________________________________________________________________

gx1010_loc_stores

int gx1010_loc_stores (int instrumentID, int operation, int storeNumber);

Purpose

Panel name:     Local Storage

Description:    Stores or Recalls the complete

state of the GX1010 to or from

local non-volatile memory.

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   Accepts the Instrument ID returned

by the Initialize function to

select the desired instrument.

Variable Type: Integer

Valid Range:   1 to  Maximum Instances of

this Instrument

The Maximum Instances of this

Instrument is the maximum number

of instruments that can be

controlled by this driver in a

single system. The value is set

internally in the driver. The

typical value is 10.

operation

Variable Type       int

Control Name:  Operation

Description:   Selects the recall or store

operation.

Variable Type: Integer

Valid Range:   0 = Recall

1 = Store

Default Value: 0 - Recall

storeNumber

Variable Type       int

Control name:   Store Number

Description:    Select the number of the store to

Recall from or Store to.

Variable type:  integer

Valid range:    0 to 9

Default value:  1

NOTE:  Store 0 is read only and contains the

default parameters as used in the *RST command.

Attempting to write to store 0 will generate an

error.

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   instrument ID out of range

-2   Operation value out of range

-3   Store Number out of range

230   Error writing to instrument

232   Instrument not initialized

240   Instrument timed out

________________________________________________________________________________

gx1010_main_offset

void gx1010_main_offset (int instrumentID, double DCOffsetV);

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   Accepts the Instrument ID returned

by the Initialize function to

select the desired instrument.

Variable Type: Integer

Valid Range:   1 to  Maximum Instances of

this Instrument

The Maximum Instances of this

Instrument is the maximum number

of instruments that can be

controlled by this driver in a

single system. The value is set

internally in the driver. The

typical value is 10.

DCOffsetV

Variable Type       double

Control Name:  DC Offset

Description:   Sets the output Offset voltage as

an EMF value expressed in Volts

Variable Type: double

Valid Range:   -10 to +10

Default Value: 0.0

________________________________________________________________________________

gx1010_main_phase

void gx1010_main_phase (int instrumentID, int phasedeg);

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   Accepts the Instrument ID returned

by the Initialize function to

select the desired instrument.

Variable Type: Integer

Valid Range:   1 to  Maximum Instances of

this Instrument

The Maximum Instances of this

Instrument is the maximum number

of instruments that can be

controlled by this driver in a

single system. The value is set

internally in the driver. The

typical value is 10.

phasedeg

Variable Type       int

Control Name:  Phase

Description:   Sets the Start/Stop Phase

parameter expressed in degrees.

Variable Type: integer

Valid Range:   -360 to +360

Default Value: 0

________________________________________________________________________________

gx1010_main_symmetry

void gx1010_main_symmetry (int instrumentID, double symmetry);

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   Accepts the Instrument ID returned

by the Initialize function to

select the desired instrument.

Variable Type: Integer

Valid Range:   1 to  Maximum Instances of

this Instrument

The Maximum Instances of this

Instrument is the maximum number

of instruments that can be

controlled by this driver in a

single system. The value is set

internally in the driver. The

typical value is 10.

symmetry

Variable Type       double

Control Name:  Symmetry

Description:   Sets the output Symmetry as a

percentage.

Variable Type: double

Valid Range:   0.1% to 99.9%

Default Value: 50.0%

Note: The symmetry can always be specified over

the range specified above but some waveform

functions impose temporary limit of their own,

see the instruction manual for more information.

________________________________________________________________________________

gx1010_main_zout

void gx1010_main_zout (int instrumentID, int zout);

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   Accepts the Instrument ID returned

by the Initialize function to

select the desired instrument.

Variable Type: Integer

Valid Range:   1 to  Maximum Instances of

this Instrument

The Maximum Instances of this

Instrument is the maximum number

of instruments that can be

controlled by this driver in a

single system. The value is set

internally in the driver. The

typical value is 10.

zout

Variable Type       int

Control Name:  Zout

Description:   Set the output impedance.

Variable Type: Integer

Valid Range:   1 = 600 Ohms

0 =  50 Ohms

Default Value: 0

________________________________________________________________________________

gx1010_new_arb

int gx1010_new_arb (int instrumentID, char fileName[]);

Purpose

Panel Name:    New ARB/STAIRCASE

Description:   Allows a new arbitary

or Staircase waveform to

be downloaded to the GX1010. After

downloading a new ARB should be

saved to ARB Local storage to

avoid it being overwritten.

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   Accepts the Instrument ID returned

by the Initialize function to

select the desired instrument.

Variable Type: Integer

Valid Range:   1 to  Maximum Instances of

this Instrument

The Maximum Instances of this

Instrument is the maximum number

of instruments that can be

controlled by this driver in a

single system. The value is set

internally in the driver. The

typical value is 10.

fileName

Variable Type       char []

Control name:   File name

Description:    This control specifies file name

from which the waveform data will

be sent to the GX1010.

Variable type:  String - max length 50 characters

Note; The file format will be as follows

ARB

SETARB followed by a space and then

1024 integers comma delimited

valid range -512 to +511

Stair

SETSTAIR followed by a space and then

1 to 16 pairs of integers length,level

comma delimited

valid range length 0 to 1024

level  -512 to +511

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   instrument ID out of range

-2   Operation value out of range

228   Error writing to instrument from file

229   Error reading from instrument to file

232   Instrument not initialized

240   Instrument timed out

________________________________________________________________________________

gx1010_send_trg

int gx1010_send_trg (int instrumentID);

Purpose

Panel name:     Send Trigger

Description:    Send a trigger command to the

GX1010. If MAN/REMOTE is selected

as the trigger source for any

active mode then that mode will be

triggered by this command. If no

modes are active then the command

will be ignored.

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   Accepts the Instrument ID returned

by the Initialize function to

select the desired instrument.

Variable Type: Integer

Valid Range:   1 to  Maximum Instances of

this Instrument

The Maximum Instances of this

Instrument is the maximum number

of instruments that can be

controlled by this driver in a

single system. The value is set

internally in the driver. The

typical value is 10.

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   instrument ID out of range

-2   Status Enable value out of range

-3   Register Value out of range

230   Error writing to instrument

232   Instrument not initialized

240   Instrument timed out

________________________________________________________________________________

gx1010_set_status

int gx1010_set_status (int instrumentID, int setStatusEnable, int value);

Purpose

Panel name:     Status Reporting

Description:    Specify the way in which status is

reported and Service Request is

generated as defined in the IEEE

488.2 standard.

The function provides access to

the three status enable registers

in the GX1010, these are:

Service Request Enable Register

- Specifies which bits, when set,

in the Status Byte register will

generate a Service Request.

Standard Event Enable Register

- Specifies which bits, when set,

in the Standard Event Status

register will set the ESB bit in

the Status Byte Register.

Parallel Poll Enable Register

- Specifies which bits, when set,

in the Status Byte register will

be used to form the ist local

message

See the IEEE 488.2 Standard and

the instrument  instruction book

for a full description of status

reporting capabilities.

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   Accepts the Instrument ID returned

by the Initialize function to

select the desired instrument.

Variable Type: Integer

Valid Range:   1 to  Maximum Instances of

this Instrument

The Maximum Instances of this

Instrument is the maximum number

of instruments that can be

controlled by this driver in a

single system. The value is set

internally in the driver. The

typical value is 10.

setStatusEnable

Variable Type       int

Control name:   Set Status Enable

Description:    Used to select which enable

register will receive the byte

value.

Variable type:  integer

Valid range:    0 = SRQ enable

1 = Standard Event Enable

2 = Parallel Poll

Default value   0 - SRQ Enable

value

Variable Type       int

Control name:   Value

Description:    A byte value which will be written

to the selected enable register.

Variable type:  integer

Valid range:    0 to 255

Default value:  0

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   instrument ID out of range

-2   Status Enable value out of range

-3   Register Value out of range

230   Error writing to instrument

232   Instrument not initialized

240   Instrument timed out

________________________________________________________________________________

gx1010_snd_cmd

int gx1010_snd_cmd (int instrumentID, char command[]);

Purpose

Panel name:     Send command string

Description:    Sends a string of bytes to the

GX1010 from a string variable.

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   Accepts the Instrument ID returned

by the Initialize function to

select the desired instrument.

Variable Type: Integer

Valid Range:   1 to  Maximum Instances of

this Instrument

The Maximum Instances of this

Instrument is the maximum number

of instruments that can be

controlled by this driver in a

single system. The value is set

internally in the driver. The

typical value is 10.

command

Variable Type       char []

Control name:   Command String

Description:    This control specifies data to be

sent to the GX1010. The absolute

maximum length allowed is 200

bytes.

Variable type:  String

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   instrument ID out of range

230   Error writing to instrument

232   Instrument not initialized

240   Instrument timed out

________________________________________________________________________________

gx1010_trig

int gx1010_trig (int instrumentID, int enable, int triggerSource,

double burstCount, double internalGenerator, int phasedeg);

Purpose

Panel Name:    Trigger (Burst)

Description:   This panel allows access to the

Trigger function of the GX1010

Parameter List

instrumentID

Variable Type       int

Control Name:  Instrument ID

Description:   Accepts the Instrument ID returned

by the Initialize function to

select the desired instrument.

Variable Type: Integer

Valid Range:   1 to  Maximum Instances of

this Instrument

The Maximum Instances of this

Instrument is the maximum number

of instruments that can be

controlled by this driver in a

single system. The value is set

internally in the driver. The

typical value is 10.

enable

Variable Type       int

Control Name:  Enable

Description:   Sets the status of Trigger.

Variable Type: Integer

Valid Range:

0 = Off

1 = On

2 = On and configure

Default Value: 0 - Off

Note: If set to ON or OFF all other controls

are ignored, to use these you must use

On + config.

triggerSource

Variable Type       int

Control Name:  Trigger Source

Description:   Sets the Trigger source.

Variable Type: Integer

Valid Range:   0 = Internal

1 = Man/Remote

2 = External

Default Value: 0 - Internal

burstCount

Variable Type       double

Control Name:  Burst Count

Description:   Sets the number of cycles of the

main generator waveform in each

burst.

Variable Type: double

Valid Range:   0.5 to 1023

Default Value: 1.0

internalGenerator

Variable Type       double

Control Name:  Internal Generator

Description:   Sets the internal trigger

frequency in Hz.

Note. This parameter is actually

the internal Trigger generator

frequency (TGEN).

Variable Type: double

Valid Range:   5e-3 to 50e+3 (5mHz to 50kHz)

Default Value: 1000

Note: The internal trigger generator period has

a setting resolution of 20us max. This

driver converts the requested frequency to a

period value and sends the resulting number to

the GX1010. This may cause a warning that the

20us resolution has been exceeded. This can

usually be ignored.

phasedeg

Variable Type       int

Control Name:  Phase

Description:   Sets the Start/Stop Phase

parameter expressed in degrees.

Variable Type: integer

Valid Range:   -360 to +360

Default Value: 0

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   instrument ID out of range

-2   Enable value out of range

-3   Trigger source value out of range

-4   Burst count value out of range

-5   Internal generator value out of range

-6   Phase value out of range

230   Error writing to instrument

232   Instrument not initialized

240   Instrument timed out