Metrix GX1025/1050 Function Generator

________________________________________________________________________________

Introduction:

This instrument driver provides programming support for Metrix GX1025/1050 Function Generator.

It contains functions for opening, configuring, taking measurements from, and

closing the instrument.

________________________________________________________________________________

Assumptions:

To successfully use this module, the following conditions must be met:

For GPIB instrument drivers:

-  the instrument is connected to the GPIB.

-  the GPIB address supplied to the initialize function must match the

GPIB address of the instrument.

For VXI instrument drivers:

-  the instrument is installed in the VXI mainframe and you are

using one of the following controller options:

Embedded controller

MXI

MXI2

GPIB-VXI

-  the logical address supplied to the initialize function must

match the logical address of the instrument.

For RS-232 instrument drivers:

-  the instrument is connected to the RS-232 interface.

-  the COM port, baud rate, parity, and timeout supplied to the

initialize function must match the settings of the instrument.

________________________________________________________________________________

Error and Status Information:

Each function in this instrument driver returns a status code that

either indicates success or describes an error or warning condition.

Your program should examine the status code from each call to an

instrument driver function to determine if an error occurred.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

The description of each instrument driver function lists possible error

codes and their meanings

________________________________________________________________________________

How To Use This Document:

Use this document as a programming reference manual.

It describes each function in the

Metrix GX1025/1050 Function Generator

instrument.  The functions appear in alphabetical order, with a description

of the function and its C syntax, a description of each parameter, and a list

of possible error codes.

________________________________________________________________________________

Function Tree Layout:

Class/Panel Name:                    Function Name:

Initialize                        gx1025_init

Waveform

Config Main                    gx1025_ConfigMain

Main Frequency                 gx1025_MainFreq

Main Amplitude                 gx1025_MainAmplitude

Main Offset                    gx1025_MainOffset

Main Symmetry                  gx1025_MainSym

Main Phase                     gx1025_MainPhase

Arb Select                     gx1025_ArbSelect

Modulation                     gx1025_setModulation

Sweep                          gx1025_setSweep

Burst                          gx1025_setBurst

Action

Main Out                       gx1025_MainOut

Trig Sweep                     gx1025_trigSweep

Trig Burst                     gx1025_trigBurst

Data

Arb Stores                     gx1025_ArbStores

Arb Recall                     gx1025_ArbRecall

Configuration

Setup                          gx1025_Setup

Send Setup                     gx1025_SendSetup

Close                             gx1025_close

________________________________________________________________________________

Metrix GX1025/1050 Function Generator

Driver for METRIX GX1025 and GX1050

________________________________________________________________________________

The following functions are in alphabetical order.

________________________________________________________________________________

gx1025_ArbRecall

int gx1025_ArbRecall (int arbNumber, int *valueArray);

Purpose

Panel Name:    Arb Recall

Description:   Recalls an arbitrary signal.

Parameter List

arbNumber

Variable Type       int

Control Name:  Arb number

Description:   Sets the number.

Variable Type: int

Valid Range:   1 to 10

Default Value: 1

valueArray

Variable Type       int (passed by reference)

Control Name:  value Array

Description:   The array recovers points of arbitrary signal.

Variable Type: long []

Size:   16384

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Unknown system error

-2   Invalid port number

-3   Port is not open

-4   Unknown IO error

-5   Unexpected internal error

-6   No serial port found

-7   Cannot open port

-11   Memory allocation error

-12   Unable to allocate system resources

-13   Invalid parameter

-14   Invalid baud rate

-24   Invalid parity mode

-34   Illegal number of data bits

-44   Illegal number of stop bits

-90   Bad file handle

-91   Error in performing file I/O

-94   Invalid count - must be 0 or greater

-97   Invalid interrupt level

-98   Default port address not defined for ports 9 through 32

-99   I/O operation timed out

-104   Break time must be between 0 and 255

-114   Invalid input queue size

-124   Invalid output queue size

-151   General I/O Error

-152   Buffer parameter is NULL

-257   No acknowledgment received after packet sent

-258   Packet not sent within retry limit

-259   Packet not received within retry limit

-260   Unexpected end of transmission received

-261   Unable to read packet number

-262   Packet number inconsistency

-263   Packet data could not be read

-264   Checksum could not be read

-265   Received checksum does not match computed checksum

-269   Invalid packet size

-300   Error opening file

-301   Error reading file

-302   Initial negative acknowledgment not received

-303   Acknowledgment not received following end of

transmission

-304   Error while writing to file

-305   Did not receive start of data nor end of

transmission when one was expected

-402   Transfer was cancelled because CAN byte was received

-503   Invalid start delay

-504   Invalid maximum number of tries

-505   Invalid wait period

-506   Invalid packet size

-507   Unable to read CRC

-508   CRC error

________________________________________________________________________________

gx1025_ArbSelect

int gx1025_ArbSelect (int channel, int arbNumber);

Purpose

Panel Name:    Arb Select

Description:   Sets up Arbitrary signal.

Parameter List

channel

Variable Type       int

Control Name:  Channel

Description:   Selects the channel.

Variable Type: Integer

Valid Range:   0 = CH1

1 = CH2

Default Value: 0 - CH1

arbNumber

Variable Type       int

Control Name:  Arb number

Description:   Select the number of Arb.

Variable Type: int

Valid Range:   2 to 10

Default Value: 2

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Unknown system error

-2   Invalid port number

-3   Port is not open

-4   Unknown IO error

-5   Unexpected internal error

-6   No serial port found

-7   Cannot open port

-11   Memory allocation error

-12   Unable to allocate system resources

-13   Invalid parameter

-14   Invalid baud rate

-24   Invalid parity mode

-34   Illegal number of data bits

-44   Illegal number of stop bits

-90   Bad file handle

-91   Error in performing file I/O

-94   Invalid count - must be 0 or greater

-97   Invalid interrupt level

-98   Default port address not defined for ports 9 through 32

-99   I/O operation timed out

-104   Break time must be between 0 and 255

-114   Invalid input queue size

-124   Invalid output queue size

-151   General I/O Error

-152   Buffer parameter is NULL

-257   No acknowledgment received after packet sent

-258   Packet not sent within retry limit

-259   Packet not received within retry limit

-260   Unexpected end of transmission received

-261   Unable to read packet number

-262   Packet number inconsistency

-263   Packet data could not be read

-264   Checksum could not be read

-265   Received checksum does not match computed checksum

-269   Invalid packet size

-300   Error opening file

-301   Error reading file

-302   Initial negative acknowledgment not received

-303   Acknowledgment not received following end of

transmission

-304   Error while writing to file

-305   Did not receive start of data nor end of

transmission when one was expected

-402   Transfer was cancelled because CAN byte was received

-503   Invalid start delay

-504   Invalid maximum number of tries

-505   Invalid wait period

-506   Invalid packet size

-507   Unable to read CRC

-508   CRC error

________________________________________________________________________________

gx1025_ArbStores

int gx1025_ArbStores (int arbNumber, char *nameStore,

double frequency_Hz, double amplitude_Vpp,

double offset_V, long valueArray16384[]);

Purpose

Panel Name:    Arb Stores

Description:   Saves an arbitrary signal.

Parameter List

arbNumber

Variable Type       int

Control Name:  Arb number

Description:   Sets the number.

Variable Type: int

Valid Range:   1 to 10

Default Value: 1

nameStore

Variable Type       char *

Control Name:  name Store

Description:   The name of arbitrary signal.

Variable Type: char *

frequency_Hz

Variable Type       double

Control Name:  Pk-Pk Output

Description:   Specifies the output level as a

peak to peak EMF value expressed

in Volts.

Variable Type: double

Valid Range:   000mV to 10.0V.

Default Value: 10V

amplitude_Vpp

Variable Type       double

Control Name:  Pk-Pk Output

Description:   Specifies the output level as a

peak to peak EMF value expressed

in Volts.

Variable Type: double

Valid Range:   000mV to 10.0V.

Default Value: 10V

offset_V

Variable Type       double

Control Name:  DC Offset

Description:   Sets the output Offset voltage as

an EMF value expressed in Volts

Variable Type: double

Valid Range:   -5.0 to +5.0

Default Value: 0.0

valueArray16384

Variable Type       long []

Control Name:  value Array

Description:   The array contains points of arbitrary signal.

Variable Type: long []

Size:   16384

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Unknown system error

-2   Invalid port number

-3   Port is not open

-4   Unknown IO error

-5   Unexpected internal error

-6   No serial port found

-7   Cannot open port

-11   Memory allocation error

-12   Unable to allocate system resources

-13   Invalid parameter

-14   Invalid baud rate

-24   Invalid parity mode

-34   Illegal number of data bits

-44   Illegal number of stop bits

-90   Bad file handle

-91   Error in performing file I/O

-94   Invalid count - must be 0 or greater

-97   Invalid interrupt level

-98   Default port address not defined for ports 9 through 32

-99   I/O operation timed out

-104   Break time must be between 0 and 255

-114   Invalid input queue size

-124   Invalid output queue size

-151   General I/O Error

-152   Buffer parameter is NULL

-257   No acknowledgment received after packet sent

-258   Packet not sent within retry limit

-259   Packet not received within retry limit

-260   Unexpected end of transmission received

-261   Unable to read packet number

-262   Packet number inconsistency

-263   Packet data could not be read

-264   Checksum could not be read

-265   Received checksum does not match computed checksum

-269   Invalid packet size

-300   Error opening file

-301   Error reading file

-302   Initial negative acknowledgment not received

-303   Acknowledgment not received following end of

transmission

-304   Error while writing to file

-305   Did not receive start of data nor end of

transmission when one was expected

-402   Transfer was cancelled because CAN byte was received

-503   Invalid start delay

-504   Invalid maximum number of tries

-505   Invalid wait period

-506   Invalid packet size

-507   Unable to read CRC

-508   CRC error

________________________________________________________________________________

gx1025_close

int gx1025_close (void);

Purpose

Panel Name:    Close

Description:   Closes the instrument and takes it

off line.

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Unknown system error

-2   Invalid port number

-3   Port is not open

-4   Unknown IO error

-5   Unexpected internal error

-6   No serial port found

-7   Cannot open port

-11   Memory allocation error

-12   Unable to allocate system resources

-13   Invalid parameter

-14   Invalid baud rate

-24   Invalid parity mode

-34   Illegal number of data bits

-44   Illegal number of stop bits

-90   Bad file handle

-91   Error in performing file I/O

-94   Invalid count - must be 0 or greater

-97   Invalid interrupt level

-98   Default port address not defined for ports 9 through 32

-99   I/O operation timed out

-104   Break time must be between 0 and 255

-114   Invalid input queue size

-124   Invalid output queue size

-151   General I/O Error

-152   Buffer parameter is NULL

-257   No acknowledgment received after packet sent

-258   Packet not sent within retry limit

-259   Packet not received within retry limit

-260   Unexpected end of transmission received

-261   Unable to read packet number

-262   Packet number inconsistency

-263   Packet data could not be read

-264   Checksum could not be read

-265   Received checksum does not match computed checksum

-269   Invalid packet size

-300   Error opening file

-301   Error reading file

-302   Initial negative acknowledgment not received

-303   Acknowledgment not received following end of

transmission

-304   Error while writing to file

-305   Did not receive start of data nor end of

transmission when one was expected

-402   Transfer was cancelled because CAN byte was received

-503   Invalid start delay

-504   Invalid maximum number of tries

-505   Invalid wait period

-506   Invalid packet size

-507   Unable to read CRC

-508   CRC error

________________________________________________________________________________

gx1025_ConfigMain

int gx1025_ConfigMain (int channel, int function);

Purpose

Panel Name:    Config Main

Description:   Sets up Function.

Parameter List

channel

Variable Type       int

Control Name:  Channel

Description:   Selects the channel.

Variable Type: Integer

Valid Range:   0 = CH1

1 = CH2

Default Value: 0 - CH1

function

Variable Type       int

Control Name:  Function

Description:   Specifies the output waveform

function.

Variable Type: Integer

Valid Range:   0 = Sine

1 = Square

2 = Ramp

3 = Pulse

4 = Noise

5 = Arb

6 = DC

Default Value: 0 - Sine

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Unknown system error

-2   Invalid port number

-3   Port is not open

-4   Unknown IO error

-5   Unexpected internal error

-6   No serial port found

-7   Cannot open port

-11   Memory allocation error

-12   Unable to allocate system resources

-13   Invalid parameter

-14   Invalid baud rate

-24   Invalid parity mode

-34   Illegal number of data bits

-44   Illegal number of stop bits

-90   Bad file handle

-91   Error in performing file I/O

-94   Invalid count - must be 0 or greater

-97   Invalid interrupt level

-98   Default port address not defined for ports 9 through 32

-99   I/O operation timed out

-104   Break time must be between 0 and 255

-114   Invalid input queue size

-124   Invalid output queue size

-151   General I/O Error

-152   Buffer parameter is NULL

-257   No acknowledgment received after packet sent

-258   Packet not sent within retry limit

-259   Packet not received within retry limit

-260   Unexpected end of transmission received

-261   Unable to read packet number

-262   Packet number inconsistency

-263   Packet data could not be read

-264   Checksum could not be read

-265   Received checksum does not match computed checksum

-269   Invalid packet size

-300   Error opening file

-301   Error reading file

-302   Initial negative acknowledgment not received

-303   Acknowledgment not received following end of

transmission

-304   Error while writing to file

-305   Did not receive start of data nor end of

transmission when one was expected

-402   Transfer was cancelled because CAN byte was received

-503   Invalid start delay

-504   Invalid maximum number of tries

-505   Invalid wait period

-506   Invalid packet size

-507   Unable to read CRC

-508   CRC error

________________________________________________________________________________

gx1025_init

int gx1025_init (void);

Purpose

Panel Name:    Initialize

Description:   Initializes the instrument in the

following ways:

- opens the instrument

- optionally performs an

identification query on the

instrument

- optionally sets the instrument

to a default configuration

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Unknown system error

-2   Invalid port number

-3   Port is not open

-4   Unknown IO error

-5   Unexpected internal error

-6   No serial port found

-7   Cannot open port

-11   Memory allocation error

-12   Unable to allocate system resources

-13   Invalid parameter

-14   Invalid baud rate

-24   Invalid parity mode

-34   Illegal number of data bits

-44   Illegal number of stop bits

-90   Bad file handle

-91   Error in performing file I/O

-94   Invalid count - must be 0 or greater

-97   Invalid interrupt level

-98   Default port address not defined for ports 9 through 32

-99   I/O operation timed out

-104   Break time must be between 0 and 255

-114   Invalid input queue size

-124   Invalid output queue size

-151   General I/O Error

-152   Buffer parameter is NULL

-257   No acknowledgment received after packet sent

-258   Packet not sent within retry limit

-259   Packet not received within retry limit

-260   Unexpected end of transmission received

-261   Unable to read packet number

-262   Packet number inconsistency

-263   Packet data could not be read

-264   Checksum could not be read

-265   Received checksum does not match computed checksum

-269   Invalid packet size

-300   Error opening file

-301   Error reading file

-302   Initial negative acknowledgment not received

-303   Acknowledgment not received following end of

transmission

-304   Error while writing to file

-305   Did not receive start of data nor end of

transmission when one was expected

-402   Transfer was cancelled because CAN byte was received

-503   Invalid start delay

-504   Invalid maximum number of tries

-505   Invalid wait period

-506   Invalid packet size

-507   Unable to read CRC

-508   CRC error

________________________________________________________________________________

gx1025_MainAmplitude

int gx1025_MainAmplitude (int channel, int function,

double amplitude_Vpp);

Purpose

Panel Name:    Main Amplitude

Description:   Sets up Function and Peak-Peak level.

Parameter List

channel

Variable Type       int

Control Name:  Channel

Description:   Selects the channel.

Variable Type: Integer

Valid Range:   0 = CH1

1 = CH2

Default Value: 0 - CH1

function

Variable Type       int

Control Name:  Function

Description:   Specifies the output waveform

function.

Variable Type: Integer

Valid Range:   0 = Sine

1 = Square

2 = Ramp

3 = Pulse

4 = Noise

5 = Arb

6 = DC

Default Value: 0 - Sine

amplitude_Vpp

Variable Type       double

Control Name:  Pk-Pk Output

Description:   Specifies the output level as a

peak to peak EMF value expressed

in Volts.

Variable Type: double

Valid Range:   000mV to 10.0V.

Default Value: 10V

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Unknown system error

-2   Invalid port number

-3   Port is not open

-4   Unknown IO error

-5   Unexpected internal error

-6   No serial port found

-7   Cannot open port

-11   Memory allocation error

-12   Unable to allocate system resources

-13   Invalid parameter

-14   Invalid baud rate

-24   Invalid parity mode

-34   Illegal number of data bits

-44   Illegal number of stop bits

-90   Bad file handle

-91   Error in performing file I/O

-94   Invalid count - must be 0 or greater

-97   Invalid interrupt level

-98   Default port address not defined for ports 9 through 32

-99   I/O operation timed out

-104   Break time must be between 0 and 255

-114   Invalid input queue size

-124   Invalid output queue size

-151   General I/O Error

-152   Buffer parameter is NULL

-257   No acknowledgment received after packet sent

-258   Packet not sent within retry limit

-259   Packet not received within retry limit

-260   Unexpected end of transmission received

-261   Unable to read packet number

-262   Packet number inconsistency

-263   Packet data could not be read

-264   Checksum could not be read

-265   Received checksum does not match computed checksum

-269   Invalid packet size

-300   Error opening file

-301   Error reading file

-302   Initial negative acknowledgment not received

-303   Acknowledgment not received following end of

transmission

-304   Error while writing to file

-305   Did not receive start of data nor end of

transmission when one was expected

-402   Transfer was cancelled because CAN byte was received

-503   Invalid start delay

-504   Invalid maximum number of tries

-505   Invalid wait period

-506   Invalid packet size

-507   Unable to read CRC

-508   CRC error

________________________________________________________________________________

gx1025_MainFreq

int gx1025_MainFreq (int channel, double frequency_Hz);

Purpose

Panel Name:    Main Frequency

Description:   Sets up Frequency.

Parameter List

channel

Variable Type       int

Control Name:  Channel

Description:   Selects the channel.

Variable Type: Integer

Valid Range:   0 = CH1

1 = CH2

Default Value: 0 - CH1

frequency_Hz

Variable Type       double

Control Name:  Frequency

Description:   The output Frequency expressed in

Hz.

Variable Type: double

Valid Range:   1e-6 to 25e+6 or 50e+6 (1Hz to 25MHz or 50MHz)

Default Value: 10kHz

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Unknown system error

-2   Invalid port number

-3   Port is not open

-4   Unknown IO error

-5   Unexpected internal error

-6   No serial port found

-7   Cannot open port

-11   Memory allocation error

-12   Unable to allocate system resources

-13   Invalid parameter

-14   Invalid baud rate

-24   Invalid parity mode

-34   Illegal number of data bits

-44   Illegal number of stop bits

-90   Bad file handle

-91   Error in performing file I/O

-94   Invalid count - must be 0 or greater

-97   Invalid interrupt level

-98   Default port address not defined for ports 9 through 32

-99   I/O operation timed out

-104   Break time must be between 0 and 255

-114   Invalid input queue size

-124   Invalid output queue size

-151   General I/O Error

-152   Buffer parameter is NULL

-257   No acknowledgment received after packet sent

-258   Packet not sent within retry limit

-259   Packet not received within retry limit

-260   Unexpected end of transmission received

-261   Unable to read packet number

-262   Packet number inconsistency

-263   Packet data could not be read

-264   Checksum could not be read

-265   Received checksum does not match computed checksum

-269   Invalid packet size

-300   Error opening file

-301   Error reading file

-302   Initial negative acknowledgment not received

-303   Acknowledgment not received following end of

transmission

-304   Error while writing to file

-305   Did not receive start of data nor end of

transmission when one was expected

-402   Transfer was cancelled because CAN byte was received

-503   Invalid start delay

-504   Invalid maximum number of tries

-505   Invalid wait period

-506   Invalid packet size

-507   Unable to read CRC

-508   CRC error

________________________________________________________________________________

gx1025_MainOffset

int gx1025_MainOffset (int channel, int function, double offset_V);

Purpose

Panel Name:    Main Offset

Description:   Sets up Function and DC Offset level.

Parameter List

channel

Variable Type       int

Control Name:  Channel

Description:   Selects the channel.

Variable Type: Integer

Valid Range:   0 = CH1

1 = CH2

Default Value: 0 - CH1

function

Variable Type       int

Control Name:  Function

Description:   Specifies the output waveform

function.

Variable Type: Integer

Valid Range:   0 = Sine

1 = Square

2 = Ramp

3 = Pulse

4 = Noise

5 = Arb

6 = DC

Default Value: 0 - Sine

offset_V

Variable Type       double

Control Name:  DC Offset

Description:   Sets the output Offset voltage as

an EMF value expressed in Volts

Variable Type: double

Valid Range:   -5.0 to +5.0

Default Value: 0.0

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Unknown system error

-2   Invalid port number

-3   Port is not open

-4   Unknown IO error

-5   Unexpected internal error

-6   No serial port found

-7   Cannot open port

-11   Memory allocation error

-12   Unable to allocate system resources

-13   Invalid parameter

-14   Invalid baud rate

-24   Invalid parity mode

-34   Illegal number of data bits

-44   Illegal number of stop bits

-90   Bad file handle

-91   Error in performing file I/O

-94   Invalid count - must be 0 or greater

-97   Invalid interrupt level

-98   Default port address not defined for ports 9 through 32

-99   I/O operation timed out

-104   Break time must be between 0 and 255

-114   Invalid input queue size

-124   Invalid output queue size

-151   General I/O Error

-152   Buffer parameter is NULL

-257   No acknowledgment received after packet sent

-258   Packet not sent within retry limit

-259   Packet not received within retry limit

-260   Unexpected end of transmission received

-261   Unable to read packet number

-262   Packet number inconsistency

-263   Packet data could not be read

-264   Checksum could not be read

-265   Received checksum does not match computed checksum

-269   Invalid packet size

-300   Error opening file

-301   Error reading file

-302   Initial negative acknowledgment not received

-303   Acknowledgment not received following end of

transmission

-304   Error while writing to file

-305   Did not receive start of data nor end of

transmission when one was expected

-402   Transfer was cancelled because CAN byte was received

-503   Invalid start delay

-504   Invalid maximum number of tries

-505   Invalid wait period

-506   Invalid packet size

-507   Unable to read CRC

-508   CRC error

________________________________________________________________________________

gx1025_MainOut

int gx1025_MainOut (int channel, int state);

Purpose

Panel Name:    Main Out

Description:   Sets up Output.

Parameter List

channel

Variable Type       int

Control Name:  Channel

Description:   Selects the channel.

Variable Type: Integer

Valid Range:   0 = CH1

1 = CH2

Default Value: 0 - CH1

state

Variable Type       int

Control Name:  Main Out

Description:   Sets the main out.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

Default Value: 0 - OFF

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Unknown system error

-2   Invalid port number

-3   Port is not open

-4   Unknown IO error

-5   Unexpected internal error

-6   No serial port found

-7   Cannot open port

-11   Memory allocation error

-12   Unable to allocate system resources

-13   Invalid parameter

-14   Invalid baud rate

-24   Invalid parity mode

-34   Illegal number of data bits

-44   Illegal number of stop bits

-90   Bad file handle

-91   Error in performing file I/O

-94   Invalid count - must be 0 or greater

-97   Invalid interrupt level

-98   Default port address not defined for ports 9 through 32

-99   I/O operation timed out

-104   Break time must be between 0 and 255

-114   Invalid input queue size

-124   Invalid output queue size

-151   General I/O Error

-152   Buffer parameter is NULL

-257   No acknowledgment received after packet sent

-258   Packet not sent within retry limit

-259   Packet not received within retry limit

-260   Unexpected end of transmission received

-261   Unable to read packet number

-262   Packet number inconsistency

-263   Packet data could not be read

-264   Checksum could not be read

-265   Received checksum does not match computed checksum

-269   Invalid packet size

-300   Error opening file

-301   Error reading file

-302   Initial negative acknowledgment not received

-303   Acknowledgment not received following end of

transmission

-304   Error while writing to file

-305   Did not receive start of data nor end of

transmission when one was expected

-402   Transfer was cancelled because CAN byte was received

-503   Invalid start delay

-504   Invalid maximum number of tries

-505   Invalid wait period

-506   Invalid packet size

-507   Unable to read CRC

-508   CRC error

________________________________________________________________________________

gx1025_MainPhase

int gx1025_MainPhase (int channel, double phase_);

Purpose

Panel Name:    Main Phase

Description:   Sets up phase.

Parameter List

channel

Variable Type       int

Control Name:  Channel

Description:   Selects the channel.

Variable Type: Integer

Valid Range:   0 = CH1

1 = CH2

Default Value: 0 - CH1

phase_

Variable Type       double

Control Name:  phase

Description:   Sets the phase between CH1 and CH2 in

degrees.

Variable Type: double

Valid Range:   0 to 360

Default Value: 0.0

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Unknown system error

-2   Invalid port number

-3   Port is not open

-4   Unknown IO error

-5   Unexpected internal error

-6   No serial port found

-7   Cannot open port

-11   Memory allocation error

-12   Unable to allocate system resources

-13   Invalid parameter

-14   Invalid baud rate

-24   Invalid parity mode

-34   Illegal number of data bits

-44   Illegal number of stop bits

-90   Bad file handle

-91   Error in performing file I/O

-94   Invalid count - must be 0 or greater

-97   Invalid interrupt level

-98   Default port address not defined for ports 9 through 32

-99   I/O operation timed out

-104   Break time must be between 0 and 255

-114   Invalid input queue size

-124   Invalid output queue size

-151   General I/O Error

-152   Buffer parameter is NULL

-257   No acknowledgment received after packet sent

-258   Packet not sent within retry limit

-259   Packet not received within retry limit

-260   Unexpected end of transmission received

-261   Unable to read packet number

-262   Packet number inconsistency

-263   Packet data could not be read

-264   Checksum could not be read

-265   Received checksum does not match computed checksum

-269   Invalid packet size

-300   Error opening file

-301   Error reading file

-302   Initial negative acknowledgment not received

-303   Acknowledgment not received following end of

transmission

-304   Error while writing to file

-305   Did not receive start of data nor end of

transmission when one was expected

-402   Transfer was cancelled because CAN byte was received

-503   Invalid start delay

-504   Invalid maximum number of tries

-505   Invalid wait period

-506   Invalid packet size

-507   Unable to read CRC

-508   CRC error

________________________________________________________________________________

gx1025_MainSym

int gx1025_MainSym (int channel, int function, double duty_);

Purpose

Panel Name:    Main Symmetry

Description:   Sets up Function and Symmetry level.

Parameter List

channel

Variable Type       int

Control Name:  Channel

Description:   Selects the channel.

Variable Type: Integer

Valid Range:   0 = CH1

1 = CH2

Default Value: 0 - CH1

function

Variable Type       int

Control Name:  Function

Description:   Specifies the output waveform

function.

Variable Type: Integer

Valid Range:   1 = Square

2 = Ramp

3 = Pulse

Default Value: 1 - Square

duty_

Variable Type       double

Control Name:  Symmetry

Description:   Sets the output Symmetry as a

percentage.

Variable Type: double

Valid Range:   20% to 80%

Default Value: 50.0%

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Unknown system error

-2   Invalid port number

-3   Port is not open

-4   Unknown IO error

-5   Unexpected internal error

-6   No serial port found

-7   Cannot open port

-11   Memory allocation error

-12   Unable to allocate system resources

-13   Invalid parameter

-14   Invalid baud rate

-24   Invalid parity mode

-34   Illegal number of data bits

-44   Illegal number of stop bits

-90   Bad file handle

-91   Error in performing file I/O

-94   Invalid count - must be 0 or greater

-97   Invalid interrupt level

-98   Default port address not defined for ports 9 through 32

-99   I/O operation timed out

-104   Break time must be between 0 and 255

-114   Invalid input queue size

-124   Invalid output queue size

-151   General I/O Error

-152   Buffer parameter is NULL

-257   No acknowledgment received after packet sent

-258   Packet not sent within retry limit

-259   Packet not received within retry limit

-260   Unexpected end of transmission received

-261   Unable to read packet number

-262   Packet number inconsistency

-263   Packet data could not be read

-264   Checksum could not be read

-265   Received checksum does not match computed checksum

-269   Invalid packet size

-300   Error opening file

-301   Error reading file

-302   Initial negative acknowledgment not received

-303   Acknowledgment not received following end of

transmission

-304   Error while writing to file

-305   Did not receive start of data nor end of

transmission when one was expected

-402   Transfer was cancelled because CAN byte was received

-503   Invalid start delay

-504   Invalid maximum number of tries

-505   Invalid wait period

-506   Invalid packet size

-507   Unable to read CRC

-508   CRC error

________________________________________________________________________________

gx1025_SendSetup

int gx1025_SendSetup (int channel, int function, double frequency_Hz,

double amplitude_Vpp, double offset_V,

double phase_, double duty_, int state);

Purpose

Panel Name:    Send Setup

Description:   Sends setup.

Parameter List

channel

Variable Type       int

Control Name:  Channel

Description:   Selects the channel.

Variable Type: Integer

Valid Range:   0 = CH1

1 = CH2

Default Value: 0 - CH1

function

Variable Type       int

Control Name:  Function

Description:   Specifies the output waveform

function.

Variable Type: Integer

Valid Range:   0 = Sine

1 = Square

2 = Ramp

3 = Pulse

4 = Noise

5 = Arb

6 = DC

Default Value: 0 - Sine

frequency_Hz

Variable Type       double

Control Name:  Frequency

Description:   The output Frequency expressed in

Hz.

Variable Type: double

Valid Range:   1e-6 to 25e+6 or 50e+6 (1Hz to 25MHz or 50MHz)

Default Value: 10kHz

amplitude_Vpp

Variable Type       double

Control Name:  Pk-Pk Output

Description:   Specifies the output level as a

peak to peak EMF value expressed

in Volts.

Variable Type: double

Valid Range:   000mV to 10.0V.

Default Value: 10V

offset_V

Variable Type       double

Control Name:  DC Offset

Description:   Sets the output Offset voltage as

an EMF value expressed in Volts

Variable Type: double

Valid Range:   -5.0 to +5.0

Default Value: 0.0

phase_

Variable Type       double

Control Name:  phase

Description:   Sets the phase between CH1 and CH2 in

degrees.

Variable Type: double

Valid Range:   0 to 360

Default Value: 0.0

duty_

Variable Type       double

Control Name:  Symmetry

Description:   Sets the output Symmetry as a

percentage.

Variable Type: double

Valid Range:   20% to 80%

Default Value: 50.0%

state

Variable Type       int

Control Name:  Main Out

Description:   Sets the main out.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

Default Value: 0 - OFF

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Unknown system error

-2   Invalid port number

-3   Port is not open

-4   Unknown IO error

-5   Unexpected internal error

-6   No serial port found

-7   Cannot open port

-11   Memory allocation error

-12   Unable to allocate system resources

-13   Invalid parameter

-14   Invalid baud rate

-24   Invalid parity mode

-34   Illegal number of data bits

-44   Illegal number of stop bits

-90   Bad file handle

-91   Error in performing file I/O

-94   Invalid count - must be 0 or greater

-97   Invalid interrupt level

-98   Default port address not defined for ports 9 through 32

-99   I/O operation timed out

-104   Break time must be between 0 and 255

-114   Invalid input queue size

-124   Invalid output queue size

-151   General I/O Error

-152   Buffer parameter is NULL

-257   No acknowledgment received after packet sent

-258   Packet not sent within retry limit

-259   Packet not received within retry limit

-260   Unexpected end of transmission received

-261   Unable to read packet number

-262   Packet number inconsistency

-263   Packet data could not be read

-264   Checksum could not be read

-265   Received checksum does not match computed checksum

-269   Invalid packet size

-300   Error opening file

-301   Error reading file

-302   Initial negative acknowledgment not received

-303   Acknowledgment not received following end of

transmission

-304   Error while writing to file

-305   Did not receive start of data nor end of

transmission when one was expected

-402   Transfer was cancelled because CAN byte was received

-503   Invalid start delay

-504   Invalid maximum number of tries

-505   Invalid wait period

-506   Invalid packet size

-507   Unable to read CRC

-508   CRC error

________________________________________________________________________________

gx1025_setBurst

int gx1025_setBurst (int channel, int type, double startPhase_,

int polarity, int source, double delay_s, int edge,

int numberOfCycles);

Purpose

Panel Name:    Burst

Description:   Sets up Burst.

Parameter List

channel

Variable Type       int

Control Name:  Channel

Description:   Selects the channel.

Variable Type: Integer

Valid Range:   0 = CH1

1 = CH2

Default Value: 0 - CH1

type

Variable Type       int

Control Name:  Type

Description:   Sets the type.

Variable Type: Integer

Valid Range:   0 = Ncycle

1 = Gated

Default Value: 0 - Ncycle

startPhase_

Variable Type       double

Control Name:  Phase

Description:   Sets the phase as degrees.

Variable Type: double

Valid Range:   0.0 to 360.0

Default Value: 0.0

polarity

Variable Type       int

Control Name:  Polarity

Description:   Sets the polarity gated.

Variable Type: Integer

Valid Range:   0 = Positive

1 = Negative

Default Value: 0 - Positive

source

Variable Type       int

Control Name:  Source

Description:   Sets the Ncycle Source.

Variable Type: Integer

Valid Range:   0 = Internal

1 = External

2 = Manual

Default Value: 0 - Internal

delay_s

Variable Type       double

Control Name:  Delay

Description:   The Delay of cycles.

Variable Type: double

Valid Range:   352ns to 1.000s

Default Value: 1.0

edge

Variable Type       int

Control Name:  Edge

Description:   Sets the Edge.

Variable Type: Integer

Valid Range:   0 = Up

1 = Down

Default Value: 0 - Up

numberOfCycles

Variable Type       int

Control Name:  Number

Description:   The number of cycles.

Variable Type: int

Valid Range:   1 to 50 000

Default Value: 1

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Unknown system error

-2   Invalid port number

-3   Port is not open

-4   Unknown IO error

-5   Unexpected internal error

-6   No serial port found

-7   Cannot open port

-11   Memory allocation error

-12   Unable to allocate system resources

-13   Invalid parameter

-14   Invalid baud rate

-24   Invalid parity mode

-34   Illegal number of data bits

-44   Illegal number of stop bits

-90   Bad file handle

-91   Error in performing file I/O

-94   Invalid count - must be 0 or greater

-97   Invalid interrupt level

-98   Default port address not defined for ports 9 through 32

-99   I/O operation timed out

-104   Break time must be between 0 and 255

-114   Invalid input queue size

-124   Invalid output queue size

-151   General I/O Error

-152   Buffer parameter is NULL

-257   No acknowledgment received after packet sent

-258   Packet not sent within retry limit

-259   Packet not received within retry limit

-260   Unexpected end of transmission received

-261   Unable to read packet number

-262   Packet number inconsistency

-263   Packet data could not be read

-264   Checksum could not be read

-265   Received checksum does not match computed checksum

-269   Invalid packet size

-300   Error opening file

-301   Error reading file

-302   Initial negative acknowledgment not received

-303   Acknowledgment not received following end of

transmission

-304   Error while writing to file

-305   Did not receive start of data nor end of

transmission when one was expected

-402   Transfer was cancelled because CAN byte was received

-503   Invalid start delay

-504   Invalid maximum number of tries

-505   Invalid wait period

-506   Invalid packet size

-507   Unable to read CRC

-508   CRC error

________________________________________________________________________________

gx1025_setModulation

int gx1025_setModulation (int channel, int function,

double ForAMAndForPM, double frequencyA_Hz,

double frequencyB_Hz, int shape, int source);

Purpose

Panel Name:    Modulation

Description:   Sets up modulation.

Parameter List

channel

Variable Type       int

Control Name:  Channel

Description:   Selects the channel.

Variable Type: Integer

Valid Range:   0 = CH1

1 = CH2

Default Value: 0 - CH1

function

Variable Type       int

Control Name:  Function

Description:   Specifies the modulation function.

Variable Type: Integer

Valid Range:   0 = OFF

1 = AM

2 = DSB_AM

3 = ASK

4 = FM

5 = FSK

6 = PM

Default Value: 0 - OFF

ForAMAndForPM

Variable Type       double

Control Name:  % for AM and  for PM modulations

Description:   Sets in percentage or degrees.

Variable Type: double

Valid Range:   0.0% to 120% for AM

0.0 to 360 for PM

Default Value: 0.0

frequencyA_Hz

Variable Type       double

Control Name:  Start Frequency

Description:   The FSK-start Frequency expressed in Hz.

Variable Type: double

Valid Range:   1e-6 to 25.0e+6 or 50.0e+6 (1Hz to 25MHz or 50MHz)

Default Value: none

frequencyB_Hz

Variable Type       double

Control Name:  Stop Frequency

Description:   The FSK-stop Frequency

expressed in Hz.

Variable Type: double

Valid Range:   1e-6 to 25.0e+6 or 50.0e+6 (1Hz to 25MHz or 50MHz)

Default Value: none

shape

Variable Type       int

Control Name:  Shape

Description:   Specifies the Shape for AM, DSB-AM, FM and PM

modulation.

Variable Type: Integer

Valid Range:   0 = SINE

1 = SQUARE

2 = TRIANGULAR

3 = UP RAMP

4 = DOWN RAMP

5 = NOISE

6 = ARB

Default Value: 0 - SINE

source

Variable Type       int

Control Name:  Source

Description:   Sets the Source.

Variable Type: Integer

Valid Range:   0 = Internal

1 = External

2 = Manual

Default Value: 0 - Internal

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Unknown system error

-2   Invalid port number

-3   Port is not open

-4   Unknown IO error

-5   Unexpected internal error

-6   No serial port found

-7   Cannot open port

-11   Memory allocation error

-12   Unable to allocate system resources

-13   Invalid parameter

-14   Invalid baud rate

-24   Invalid parity mode

-34   Illegal number of data bits

-44   Illegal number of stop bits

-90   Bad file handle

-91   Error in performing file I/O

-94   Invalid count - must be 0 or greater

-97   Invalid interrupt level

-98   Default port address not defined for ports 9 through 32

-99   I/O operation timed out

-104   Break time must be between 0 and 255

-114   Invalid input queue size

-124   Invalid output queue size

-151   General I/O Error

-152   Buffer parameter is NULL

-257   No acknowledgment received after packet sent

-258   Packet not sent within retry limit

-259   Packet not received within retry limit

-260   Unexpected end of transmission received

-261   Unable to read packet number

-262   Packet number inconsistency

-263   Packet data could not be read

-264   Checksum could not be read

-265   Received checksum does not match computed checksum

-269   Invalid packet size

-300   Error opening file

-301   Error reading file

-302   Initial negative acknowledgment not received

-303   Acknowledgment not received following end of

transmission

-304   Error while writing to file

-305   Did not receive start of data nor end of

transmission when one was expected

-402   Transfer was cancelled because CAN byte was received

-503   Invalid start delay

-504   Invalid maximum number of tries

-505   Invalid wait period

-506   Invalid packet size

-507   Unable to read CRC

-508   CRC error

________________________________________________________________________________

gx1025_setSweep

int gx1025_setSweep (int channel, double startFrequency_Hz,

double stopFrequency_Hz, double time_s, int source,

int trig, int law, int direction);

Purpose

Panel Name:    Sweep

Description:   Sets up Sweep.

Parameter List

channel

Variable Type       int

Control Name:  Channel

Description:   Selects the channel.

Variable Type: Integer

Valid Range:   0 = CH1

1 = CH2

Default Value: 0 - CH1

startFrequency_Hz

Variable Type       double

Control Name:  Start Frequency

Description:   The Sweep-start Frequency expressed in Hz.

Variable Type: double

Valid Range:   0.001 to 25.0e+6 or 50.0e+6(1mHz to 25MHz or 50MHz)

Default Value: 1000

stopFrequency_Hz

Variable Type       double

Control Name:  Stop Frequency

Description:   The Sweep-stop Frequency expressed in Hz.

Variable Type: double

Valid Range:   0.001 to 25.0e+6 or 50.0e+6(1mHz to 25MHz or 50MHz)

Default Value: 1000

time_s

Variable Type       double

Control Name:  Time

Description:   The Sweep-time expressed in s

Variable Type: double

Valid Range:   0.001 to 500.0 (1ms to 500s)

Default Value: 1

source

Variable Type       int

Control Name:  Source

Description:   Sets the Source.

Variable Type: Integer

Valid Range:   0 = Internal

1 = External

2 = Manual

Default Value: 0 - Internal

trig

Variable Type       int

Control Name:  Trig

Description:   Sets the Trig Out.

Variable Type: Integer

Valid Range:   0 = Off

1 = On

Default Value: 0 - Off

law

Variable Type       int

Control Name:  Law

Description:   Sets the law.

Variable Type: Integer

Valid Range:   0 = Line

1 = Log

Default Value: 0 - Line

direction

Variable Type       int

Control Name:  Direction

Description:   Sets the direction.

Variable Type: Integer

Valid Range:   0 = Up

1 = Down

Default Value: 0 - Up

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Unknown system error

-2   Invalid port number

-3   Port is not open

-4   Unknown IO error

-5   Unexpected internal error

-6   No serial port found

-7   Cannot open port

-11   Memory allocation error

-12   Unable to allocate system resources

-13   Invalid parameter

-14   Invalid baud rate

-24   Invalid parity mode

-34   Illegal number of data bits

-44   Illegal number of stop bits

-90   Bad file handle

-91   Error in performing file I/O

-94   Invalid count - must be 0 or greater

-97   Invalid interrupt level

-98   Default port address not defined for ports 9 through 32

-99   I/O operation timed out

-104   Break time must be between 0 and 255

-114   Invalid input queue size

-124   Invalid output queue size

-151   General I/O Error

-152   Buffer parameter is NULL

-257   No acknowledgment received after packet sent

-258   Packet not sent within retry limit

-259   Packet not received within retry limit

-260   Unexpected end of transmission received

-261   Unable to read packet number

-262   Packet number inconsistency

-263   Packet data could not be read

-264   Checksum could not be read

-265   Received checksum does not match computed checksum

-269   Invalid packet size

-300   Error opening file

-301   Error reading file

-302   Initial negative acknowledgment not received

-303   Acknowledgment not received following end of

transmission

-304   Error while writing to file

-305   Did not receive start of data nor end of

transmission when one was expected

-402   Transfer was cancelled because CAN byte was received

-503   Invalid start delay

-504   Invalid maximum number of tries

-505   Invalid wait period

-506   Invalid packet size

-507   Unable to read CRC

-508   CRC error

________________________________________________________________________________

gx1025_Setup

int gx1025_Setup (int channel, int *function, double *frequency_Hz,

double *amplitude_Vpp, double *offset_V,

double *phase_, double *duty_, int *state);

Purpose

Panel Name:    Setup

Description:   Reads setup.

Parameter List

channel

Variable Type       int

Control Name:  Channel

Description:   Selects the channel.

Variable Type: Integer

Valid Range:   0 = CH1

1 = CH2

Default Value: 0 - CH1

function

Variable Type       int (passed by reference)

Control Name:  Function

Description:   Returns the function.

Variable Type: Integer

Valid Range:   0 = Sine

1 = Square

2 = Ramp

3 = Pulse

4 = Noise

5 = Arb

6 = DC

frequency_Hz

Variable Type       double (passed by reference)

Control Name:  Frequency

Description:   Returns the frequency.

Variable Type: Double

Valid Range:   1e-6 to 25e+6 or 50e+6 (1Hz to 25MHz or 50MHz)

amplitude_Vpp

Variable Type       double (passed by reference)

Control Name:  Amplitude

Description:   Returns the amplitude.

Variable Type: Double

Valid Range:   000mV to 10.0V.

offset_V

Variable Type       double (passed by reference)

Control Name:  Offset

Description:   Returns the offset.

Variable Type: Double

Valid Range:   -5.0 to +5.0

phase_

Variable Type       double (passed by reference)

Control Name:  Phase

Description:   Returns the phase.

Variable Type: Double

Valid Range:   0 to 360

duty_

Variable Type       double (passed by reference)

Control Name:  Duty

Description:   Returns the duty.

Variable Type: Double

Valid Range:   20% to 80%

state

Variable Type       int (passed by reference)

Control Name:  State

Description:   Returns the state.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Unknown system error

-2   Invalid port number

-3   Port is not open

-4   Unknown IO error

-5   Unexpected internal error

-6   No serial port found

-7   Cannot open port

-11   Memory allocation error

-12   Unable to allocate system resources

-13   Invalid parameter

-14   Invalid baud rate

-24   Invalid parity mode

-34   Illegal number of data bits

-44   Illegal number of stop bits

-90   Bad file handle

-91   Error in performing file I/O

-94   Invalid count - must be 0 or greater

-97   Invalid interrupt level

-98   Default port address not defined for ports 9 through 32

-99   I/O operation timed out

-104   Break time must be between 0 and 255

-114   Invalid input queue size

-124   Invalid output queue size

-151   General I/O Error

-152   Buffer parameter is NULL

-257   No acknowledgment received after packet sent

-258   Packet not sent within retry limit

-259   Packet not received within retry limit

-260   Unexpected end of transmission received

-261   Unable to read packet number

-262   Packet number inconsistency

-263   Packet data could not be read

-264   Checksum could not be read

-265   Received checksum does not match computed checksum

-269   Invalid packet size

-300   Error opening file

-301   Error reading file

-302   Initial negative acknowledgment not received

-303   Acknowledgment not received following end of

transmission

-304   Error while writing to file

-305   Did not receive start of data nor end of

transmission when one was expected

-402   Transfer was cancelled because CAN byte was received

-503   Invalid start delay

-504   Invalid maximum number of tries

-505   Invalid wait period

-506   Invalid packet size

-507   Unable to read CRC

-508   CRC error

________________________________________________________________________________

gx1025_trigBurst

int gx1025_trigBurst (int channel);

Purpose

Panel Name:    Trig Burst

Description:   Arms trigger of burst.

Parameter List

channel

Variable Type       int

Control Name:  Channel

Description:   Selects the channel.

Variable Type: Integer

Valid Range:   0 = CH1

1 = CH2

Default Value: 0 - CH1

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Unknown system error

-2   Invalid port number

-3   Port is not open

-4   Unknown IO error

-5   Unexpected internal error

-6   No serial port found

-7   Cannot open port

-11   Memory allocation error

-12   Unable to allocate system resources

-13   Invalid parameter

-14   Invalid baud rate

-24   Invalid parity mode

-34   Illegal number of data bits

-44   Illegal number of stop bits

-90   Bad file handle

-91   Error in performing file I/O

-94   Invalid count - must be 0 or greater

-97   Invalid interrupt level

-98   Default port address not defined for ports 9 through 32

-99   I/O operation timed out

-104   Break time must be between 0 and 255

-114   Invalid input queue size

-124   Invalid output queue size

-151   General I/O Error

-152   Buffer parameter is NULL

-257   No acknowledgment received after packet sent

-258   Packet not sent within retry limit

-259   Packet not received within retry limit

-260   Unexpected end of transmission received

-261   Unable to read packet number

-262   Packet number inconsistency

-263   Packet data could not be read

-264   Checksum could not be read

-265   Received checksum does not match computed checksum

-269   Invalid packet size

-300   Error opening file

-301   Error reading file

-302   Initial negative acknowledgment not received

-303   Acknowledgment not received following end of

transmission

-304   Error while writing to file

-305   Did not receive start of data nor end of

transmission when one was expected

-402   Transfer was cancelled because CAN byte was received

-503   Invalid start delay

-504   Invalid maximum number of tries

-505   Invalid wait period

-506   Invalid packet size

-507   Unable to read CRC

-508   CRC error

________________________________________________________________________________

gx1025_trigSweep

int gx1025_trigSweep (int channel);

Purpose

Panel Name:    Trig Sweep

Description:   Arms trigger of sweep.

Parameter List

channel

Variable Type       int

Control Name:  Channel

Description:   Selects the channel.

Variable Type: Integer

Valid Range:   0 = CH1

1 = CH2

Default Value: 0 - CH1

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Unknown system error

-2   Invalid port number

-3   Port is not open

-4   Unknown IO error

-5   Unexpected internal error

-6   No serial port found

-7   Cannot open port

-11   Memory allocation error

-12   Unable to allocate system resources

-13   Invalid parameter

-14   Invalid baud rate

-24   Invalid parity mode

-34   Illegal number of data bits

-44   Illegal number of stop bits

-90   Bad file handle

-91   Error in performing file I/O

-94   Invalid count - must be 0 or greater

-97   Invalid interrupt level

-98   Default port address not defined for ports 9 through 32

-99   I/O operation timed out

-104   Break time must be between 0 and 255

-114   Invalid input queue size

-124   Invalid output queue size

-151   General I/O Error

-152   Buffer parameter is NULL

-257   No acknowledgment received after packet sent

-258   Packet not sent within retry limit

-259   Packet not received within retry limit

-260   Unexpected end of transmission received

-261   Unable to read packet number

-262   Packet number inconsistency

-263   Packet data could not be read

-264   Checksum could not be read

-265   Received checksum does not match computed checksum

-269   Invalid packet size

-300   Error opening file

-301   Error reading file

-302   Initial negative acknowledgment not received

-303   Acknowledgment not received following end of

transmission

-304   Error while writing to file

-305   Did not receive start of data nor end of

transmission when one was expected

-402   Transfer was cancelled because CAN byte was received

-503   Invalid start delay

-504   Invalid maximum number of tries

-505   Invalid wait period

-506   Invalid packet size

-507   Unable to read CRC

-508   CRC error